CONNECTICUT  DEPARTMENT  OF  PUBLIC  HEALTH

DIVISION  OF  HEALTH  SYSTEMS  REGULATION

410 Capitol Ave., M.S.# 12HSR, Hartford, CT 06134;  860-509-7400

VERTEBRATE   ANIMAL   BIOSAFETY  INSPECTION  CHECKLIST.

Laboratory: ___________________________________    Registration No.: __________

Address:     _______________________________________________________________

Director:     ____________________________________         Room No.: ____________

Inspector:   _______________________________    Date of Inspection: ____________

Agents Used:	  Bacterial	  Viral 		  Parasitic	  Fungal

Transmissible spongiform encephalopathies	  Rickettsial

Biosafety Level:   2      3 ;    Animals Used:   Mice;      Rats;

Other (specify:______________ .)

A. Standard Microbiological Practices: ANIMAL Biosafety Level 2 Yes No NA

Animal Biosafety Level 2 involves practices for work with those agents associated with human disease. It addresses hazards from ingestion as well as from percutaneous and mucous membrane exposure. ABSL-2 builds upon the practices, procedures, containment equipment, and facility requirements of ABSL-1.

1. The animal facility director establishes policies, procedures, and protocols for emergency situations. Each project is subject to pre-approval by the Institutional Animal Care and Use Committee (IACUC) and the Institutional Biohazard Committee (IBC). Any special practices are approved at this time. Aside from the standard policies, procedures, and protocols for emergency situations established by the facility director, appropriate special policies and procedures should be developed as needed and approved by the Institutional Animal Care and Use Committee (IACUC) and the Institutional Biohazard Committee (IBC).

2. Access to the animal room is limited to the fewest number of individuals possible. Personnel who must enter the room for program or service purposes when work is in progress are advised of the potential hazard.

3. An appropriate medical surveillance program is in place. All personnel receive appropriate immunizations or tests for the agents handled or potentially present (e.g., hepatitis B vaccine, TB skin testing). When appropriate, a serum surveillance system should be implemented.

4. A biosafety manual is prepared or adopted. Personnel are advised of special hazards, and are required to read and follow instructions on practices and procedures.

5. Eating, drinking, smoking, handling contact lenses, applying cosmetics, and storing food for human use should only be done in designated areas and are not permitted in animal or procedure rooms.

6. All procedures are carefully performed to minimize the creation of aerosols or splatters.

7. Equipment and work surfaces in the room are routinely decontaminated with an effective disinfectant after work with the infectious agent, and especially after overt spills, splashes, or other contamination by infectious materials.

Page  1                                                              Rev.  7/25/01

A. Standard Microbiological Practices: ANIMAL Biosafety Level 2 Yes No NA

8. All infectious samples are collected, labeled, transported, and processed in a manner that contains and prevents transmission of the agent(s). All wastes from the animal room (including animal tissues, carcasses, contaminated bedding, unused feed, sharps, and other refuse) are transported from the animal room in leak-proof, covered containers for appropriate disposal in compliance with applicable institutional or local requirements. The outer surface of the containers is disinfected prior to moving the material. Autoclaving of the contents prior to incineration is recommended.

9. Policies for the safe handling of sharps are instituted: Needles and syringes or other sharp instruments are restricted for use in the animal facility only when there is no alternative, such as for parenteral injection, blood collection, or aspiration of fluids from laboratory animals and diaphragm bottles. Syringes that re-sheathe the needle, needle-less systems, and other safe devices should be used when appropriate. c. Plasticware should be substituted for glassware whenever possible.

10. Personnel wash their hands after handling cultures and animals, after removing gloves, and before leaving the animal facility.

11. A biohazard sign must be posted on the entrance to the animal room whenever infectious agents are present. The hazard warning sign identifies the infectious agent(s) in use, lists the name and telephone number of the responsible person(s), and indicates the special requirements (e.g., the need for immunizations and respirators) for entering the animal room.

12. An insect and rodent control program is in effect (see Appendix G).

B. Special Practices ABSL- 2 Yes No NA

Animal care laboratory and support personnel receive appropriate training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures. Personnel receive annual updates, or additional training as necessary for procedural or policy changes. Records of all training provided are maintained. In general, persons who may be at increased risk of acquiring infection, or for whom infection might be unusually hazardous, are not allowed in the animal facility unless special procedures can eliminate the extra risk.

2. Only animals used for the experiment(s) are allowed in the room.

All equipment must be appropriately decontaminated prior to removal from the room.

Spills and accidents which result in overt exposures to infectious materials must be immediately reported to the facility director. Medical evaluation, surveillance, and treatment are provided as appropriate and written records are maintained.

Page  2.

C. Safety Equipment (Primary Barriers) ABSL-2 Yes No NA

Gowns, uniforms, or laboratory coats are worn while in the animal room. The laboratory coat is removed and left in the animal room. Gowns, uniforms, and laboratory coats are removed before leaving the animal facility. Gloves are worn when handling infected animals and when skin contact with infectious materials is unavoidable.

Personal protective equipment is used based on risk assessment determinations. Appropriate face/eye and respiratory protection is worn by all personnel entering animal rooms that house nonhuman primates.

Biological safety cabinets, other physical containment devices, and/or personal protective equipment (e.g., respirators, face shields) are used whenever conducting procedures with a high potential for creating aerosols. These include necropsy of infected animals, harvesting of tissues or fluids from infected animals or eggs, or intranasal inoculation of animals.

4. When needed, animals are housed in primary biosafety containment equipment appropriate for the animal species. Filter top cages are always handled in properly designed and operating animal biocontainment cabinets recommended for rodents.

D. Laboratory Facilities (Secondary Barriers) ABSL-2 Yes No NA

The animal facility is separated from areas that are open to unrestricted Personnel traffic within the building.

Access to the facility is limited by secure locked doors. External doors are self-closing and self-locking. Doors to animal rooms open inward, are self-closing, and are kept closed when experimental animals are present. Cubicle room inner doors may open outward or be horizontal or vertical sliding.

3. The animal facility is designed, constructed, and maintained to facilitate cleaning and housekeeping. The interior surfaces (walls, floors, and ceilings) are water resistant.

4. Internal facility appurtenances, such as light fixtures, air ducts, and utility pipes, are arranged to minimize horizontal surface areas.

5. Any windows must be resistant to breakage. Where possible, windows should be sealed. If the animal facility has windows that open, they are fitted with fly screens.

6. If floor drains are provided, the traps are always filled with an appropriate disinfectant.

Exhaust air is discharged to the outside without being recirculated to other rooms. Ventilation should be provided in accordance with criteria from Guide for Care and Use of Laboratory Animals, latest edition. The direction of airflow in the animal facility is inward; animal rooms should maintain negative pressure compared to adjoining hallways.

Cages are washed manually or in an appropriate cage washer. The mechanical cage washer should have a final rinse temperature of at least 180F.

An autoclave is available in the animal facility to decontaminate infectious waste.

A hand washing sink is in the animal room where infected animals are housed, as well as elsewhere in the facility.

Illumination is adequate for all activities, avoiding reflections and glare that could impede vision.

Page  3.

A. Standard Microbiological Practices: (ABSL-3) Yes No NA

Animal Biosafety Level 3 involves practices suitable for work with animals infected with indigenous or exotic agents that present the potential of aerosol transmission and of causing serious or potentially lethal disease. ABSL-3 builds upon the standard practices, procedures, containment equipment, and facility requirements of ABSL-2.

1. Aside from the standard policies, procedures, and protocols for emergency situations established by the facility director, appropriate special policies and procedures should be developed as needed and approved by the Institutional Animal Care and Use Committee (IACUC) and the Institutional Biosafety Committee (IBC).

2. The laboratory or animal facility director limits access to the animal room to the fewest number of individuals possible. Personnel who must enter the room for program or service purposes when work is in progress are advised of the potential hazard.

3. An appropriate medical surveillance program is in place. All personnel receive appropriate immunizations or tests for the agents handled or potentially present (e.g., hepatitis B vaccine, TB skin testing). When appropriate, a serum surveillance system should be implemented. In general, persons who may be at increased risk of acquiring infection, or for whom infection might have serious consequences, are not allowed in the animal facility unless special procedures can eliminate the extra risk. Assessment should be made by the occupational health physician.

4. A biosafety manual is prepared or adopted. Personnel are advised of special hazards, and are required to read and follow instructions on practices and procedures.

5. Eating, drinking, smoking, handling contact lenses, applying cosmetics, and storing food for human use should be done only in designated areas and are not permitted in animal or procedure rooms.

6. All procedures are carefully performed to minimize the creation of aerosols or splatters.

7. Equipment and work surfaces in the room are routinely decontaminated with an effective disinfectant after work with the infectious agent, and especially after overt spills, splashes, or other contamination by infectious materials.

8. All wastes from the animal room (including animal tissues, carcasses, contaminated bedding, unused feed, sharps, and other refuse animal tissues) are transported from the animal room in leak-proof, covered containers for appropriate disposal in compliance with applicable institutional or local requirements. Incineration is recommended. The outer surface of the containers is disinfected prior to moving the material (see Special Practices #3 below).

9. Policies for the safe handling of sharps are instituted. Needles and syringes or other sharp instruments are restricted in the animal facility for use only when there is no alternative, such as for parenteral injection, blood collection, or aspiration of fluids from laboratory animals and diaphragm bottles. Syringes that re-sheathe the needle, needle-less systems, and other safe devices should be used when appropriate. c. Plasticware should be substituted for glassware whenever possible.

10. Personnel wash their hands after handling cultures and animals, after removing gloves, and before leaving the animal facility.

Page 4

A. Standard Microbiological Practices: (ABSL-3) Yes No NA

11. A biohazard sign must be posted on the entrance to the animal room whenever infectious agents are present. The hazard warning sign identifies the infectious agent(s) in use, lists the name and telephone number of the responsible person(s), and indicates the special requirements for entering the animal room (e.g., the need for immunizations and respirators).

12. All infectious samples are collected, labeled, transported, and processed in a manner that contains and prevents transmission of the agent(s).

13. Laboratory and support personnel receive appropriate training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures. As necessary, personnel receive updates and/or additional training on procedural or policy changes. Records of all training provided are maintained.

14. An insect and rodent control program is in effect.

B. Special Practices (ABSL-3) Yes No NA

Cages are autoclaved or thoroughly decontaminated before bedding is removed and before they are cleaned and washed. Equipment must be decontaminated according to any local, state, or federal regulations before being packaged for transport or removal from the facility for repair or maintenance.

2. A spill procedure is developed and posted. Only personnel properly trained and equipped to work with infectious materials are to clean up spills. Spills and accidents that result in overt exposures to infectious materials must be immediately reported to the facility director. Medical evaluation, surveillance, and treatment are provided as appropriate and written records are maintained.

3. All wastes from the animal room must be autoclaved prior to incineration or other appropriate terminal treatment.

4. Materials not related to the experiment (e.g., plants, animals) are not permitted in the animal room.

C. Safety Equipment (Primary Barriers) (ABSL-3) Yes No NA

Uniforms or scrub suits are worn by personnel entering the animal room. Wrap-around or solid-front gowns should be worn over this clothing. Front-button laboratory coats are unsuitable. The gown must be removed and left in the animal room. Before leaving the animal facility, scrub suits and uniforms are removed and appropriately contained and decontaminated prior to laundering or disposal.

2. Personal protective equipment used is based on risk assessment determinations. Personal protective equipment is used for all activities involving manipulations of infectious material or infected animals. Personnel wear gloves when handling infected animals. Gloves are removed aseptically and autoclaved with other animal room wastes before disposal. c. Appropriate face/eye and respiratory protection (e.g., respirators and face shields) is worn by all personnel entering animal rooms. d. Boots, shoe covers, or other protective footwear, and disinfectant foot baths are available and used where indicated.

Page 5

C. Safety Equipment (Primary Barriers) (ABSL-3) Yes No NA

3. The risk of infectious aerosols from infected animals or their bedding also can be reduced if animals are housed in containment caging systems, such as open cages placed in inward flow ventilated enclosures (e.g., laminar flow cabinets), solid wall and bottom cages covered with filter bonnets, or other equivalent primary containment systems.

4. Biological safety cabinets and other physical containment devices are used whenever conducting procedures with a potential for creating aerosols. These include necropsy of infected animals, harvesting of tissues or fluids from infected animals or eggs, or intranasal inoculation of animals. At BSL-3, all work should be done in a primary barrier; otherwise respirators should be worn by personnel in the room.

D. Laboratory Facilities (Secondary Barriers) (ABSL-3) Yes No NA

The animal facility is separated from areas that are open to unrestricted personnel traffic within the building.

Access to the facility is limited by a self-closing and self-locking door. This exterior entry door may be controlled by a key lock, card key, or proximity reader. Entry into the animal room is via a double-door entry which includes a change room and shower(s). An additional double-door access (air-lock) or double-doored autoclave may be provided for movement of supplies and wastes into and out of the facility, respectively. Doors to animal rooms open inward and are self-closing. Doors to cubicles inside an animal room may open outward or slide horizontally or vertically.

The animal facility is designed, constructed, and maintained to facilitate cleaning and housekeeping. The interior surfaces (walls, floors, and ceilings) are water resistant. Penetrations in floors, walls and ceiling surfaces are sealed and openings around ducts and the spaces between doors and frames are capable of being sealed to facilitate decontamination.

A hands-free or automatically operated hand washing sink is provided in each animal room near the exit door. The sink trap is filled with an appropriate disinfectant after each use.

Internal facility appurtenances, such as light fixtures, air ducts, and utility pipes, are arranged to minimize horizontal surface areas.

Windows are not recommended. Any windows must be resistant to breakage and must be sealed.

7. If floor drains are provided, they are always filled with an appropriate disinfectant.

8. Ventilation should be provided in accordance with criteria from the Guide for Care and Use of Laboratory Animals, latest edition. A ducted exhaust air ventilation system is provided. This system creates directional airflow which draws air into the laboratory from "clean" areas and toward "contaminated" areas. The exhaust air is not recirculated to any other area of the building. Filtration and other treatments of the exhaust air may not be required, but should be considered based on site requirements, and specific agent manipulations and use conditions. The exhaust must be dispersed away from occupied areas and air intakes, or the exhaust must be HEPA- filtered. Personnel must verify that the direction of the airflow (into the animal areas) is proper. It is recommended that a visual monitoring device that indicates and confirms directional inward airflow be provided at the animal room entry. Consideration should be given to installing an HVAC control system to prevent sustained positive pressurization of the animal spaces. Audible alarms should be considered to notify personnel of HVAC system failure.

Page 6

D. Laboratory Facilities (Secondary Barriers) (ABSL-3) Yes No NA

9. HEPA-filtered exhaust air from a Class II biological safety cabinet can be recirculated into the animal room if the cabinet is tested and certified at least annually. When exhaust air from Class II safety cabinets is to be discharged to the outside through the building exhaust air system, the cabinets must be connected in a manner that avoids any interference with the air balance of the cabinets or the building exhaust system (e.g., an air gap between the cabinet exhaust and the exhaust duct). When Class III biological safety cabinets are used, they should be directly connected to the exhaust system. If the Class III cabinets are connected to the supply system, it is done in a manner that prevents positive pressurization of the cabinets (see Appendix A).

Cages are washed in a cage washer. The mechanical cage washer has a final rinse temperature of at least 180F.

11. An autoclave is available which is convenient to the animal rooms where the biohazard is contained. The autoclave is utilized to decontaminate infectious waste before moving it to other areas of the facility.

12. If vacuum service (i.e., central or local) is provided, each service connection should be fitted with liquid disinfectant traps and an in-line HEPA filter, placed as near as practicable to each use point or service cock. Filters are installed to permit in-place decontamination and replacement.

13. Illumination is adequate for all activities, avoiding reflections and glare that could impede vision.

14. The completed Biosafety Level 3 facility design and operational procedures must be documented. The facility must be tested for verification that the design and operational parameters have been met prior to operation. Facilities should be re-verified at least annually against these procedures as modified by operational experience.

15. Additional environmental protection (e.g., personnel showers, HEPA filtration of exhaust air, containment of other piped services, and the provision of effluent decontamination) should be considered if recommended by the agent summary statement, as determined by risk assessment of the site conditions, or other applicable federal, state, or local regulations.

E. Connecticut Standard Requirements; ABSL 3 Yes No NA

1. A biosafety committee reviews safety procedures and health issues pursuant to: CDC/NIH Guideline: Biosafety in Microbiological and Biomedical Laboratories, 4 rd Edition, May 1999; and NIH Guidelines For Research Involving Recombinant DNA Molecules Jan. 2001. 2. The biosafety committee of a biolevel 3 laboratory meets at least annually with a representative of the Connecticut Department of Public Health. 3. The biolevel 3 laboratory safety committee meeting minutes are forwarded to the Connecticut Department of Public Health. 4. The institution which operates a biosafety level 3 laboratory reports any sanctions imposed on the laboratory by the NIH, CDC, Dept. of Defense or any other government agency or institution, for incidents occurring at the laboratory.

Page 7

F. Connecticut Standard Requirements; Registered Laboratories Yes No NA

Chemical Safety:

1. The laboratory has a Chemical Hygiene Plan readily available to employees.

2. Employees are provided training to ensure that they are aware of hazards.

3. Written emergency procedures are available in case of spills and accidents.

4. The chemical hygiene officer inspects storage facilities to ensure that incompatible reagents are not stored together.

5. Compressed gas cylinders are secured.

Safety Committee: Yes No NA

1. As specified in CT General Statutes, Sec. 31-40v, each employer of 25 or more employees shall administer a safety and health committee.

Biosafety Level 1 represents a basic level of containment that relies on standard microbiological practices with no special primary or secondary barriers recommended, other than a sink for handwashing. (1)

Biosafety Level 2 practices, equipment, and facility design and construction are applicable to clinical, diagnostic, teaching, and other laboratories in which work is done with a broad spectrum of indigenous moderate-risk agents that are present in the community and associated with human disease of varying severity. Biosafety Level 2 is appropriate when work is done with any human-derived blood, body fluids, tissues, or primary human cell lines where the presence of an infectious agent may be unknown. Laboratory personnel working with human-derived materials should refer to the OSHA Bloodborne Pathogen Standard for specific requirements. (1)

Biosafety Level 3 is applicable to clinical, diagnostic, teaching, research, or production facilities in which work is done with indigenous or exotic agents which may cause serious or potentially lethal disease as a result of exposure by the inhalation route. Laboratory personnel have specific training in handling pathogenic and potentially lethal agents, and are supervised by competent scientists who are experienced in working with these agents. All procedures involving the manipulation of infectious materials are conducted within biological safety cabinets or other physical containment devices, or by personnel wearing appropriate personal protective clothing and equipment. The laboratory has special engineering and design features. It is recognized, however, that some existing facilities may not have all the facility features recommended for Biosafety Level 3 (i.e., double-door access zone and sealed penetrations). In this circumstance, an acceptable level of safety for the conduct of routine procedures, (e.g., diagnostic procedures involving the propagation of an agent for identification, typing, susceptibility testing, etc.), may be achieved in a Biosafety Level 2 facility, providing 1) the exhaust air from the laboratory room is discharged to the outdoors, 2) the ventilation to the laboratory is balanced to provide directional airflow into the room, 3) access to the laboratory is restricted when work is in progress, and 4) the recommended Standard Microbiological Practices, Special Practices, and Safety Equipment for Biosafety Level 3 are rigorously followed. The decision to implement this modification of Biosafety Level 3 recommendations should be made only by the laboratory director.(1)

Intended use: This inspection checklist is intended to be used for evaluating the biosafety standards utilized in microbiology laboratories. A performance based approach must be taken when evaluating compliance with the requirements and recommendations contained in this checklist. Laboratories may adopt new technologies and innovative techniques to meet the intent of the standards, providing that adequate safety precautions are enforced to protect public health by ensuring containment of microorganisms.

Personnel  Contacted: _________________________________     Telephone No.: _________________

Page   8.

Biological  Safety  Cabinets  (BSC):

Class I 75 FPM; Personnel / Env. Protection / Not Product Protection; 100% HEPA filtered exhaust to room

Class  II, Type A 75 fpm opening. HEPA air is 70% Recirculated to inside cabinet / 30 % Exhaust to room

Class II, Type B1  100 fpm; HEPA Air from above, classic design; 30% Recirculated / 70 % Bldg.Exhaust

Class II, Type B2  100 fpm; HEPA air  above;  Radionuclides / Toxic Chemicals;   100 % Bldg. Exhaust

Class II, Type B3  100 fpm opening, front grill, ducted to rear side; 70% Recirculated / 30 % Bldg. exhaust

Class III  fpm is N/A, supply & exhaust is double filtered; Total Containment - Gas Tight - Glove Port –

For BL-3/4; 70% Recirculated / 30 % Bldg. Exhaust. May have dunk tank or pass-through box to autoclave.

* * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * * *

# References

1.  	Biosafety  in  Microbiological  and  Biomedical  Laboratories,   4th Edition,  May 1999.

US Govt. Printing Office.  Tel: 202-512-1800;  HHS Pub. No. (CDC) 93-8395;  Cost:  $12.00

Internet Copy Available Free  At:  www.cdc.gov

- Guide for the Care and Use of Laboratory Animals,  National Academy Press, Washington, DC, 1996.

3.  U.S. Dept. of Agriculture.  Laboratory Animal Welfare Regulations – 9 CFR, Subchapter A, Parts 1, 2, 3.

4.  Guidelines  For  Research  Involving  Recombinant  DNA  Molecules  (NIH  Guidelines),

January 2001.  Federal Register Dec. 12 & 28, 2000.  (66 FR 1146).  Available at  www.absa.org

5.  Connecticut  Public  Health Code  Regulations,  Sections  19a-36-A25  to  19a-36-A31.

Registration of laboratories which utilize any living agent capable of causing human infections or reportable diseases of

man, or which are used to secure evidence bearing upon the presence or absence of such living agents or the illness caused.

6.   Occupational  Exposure to  Bloodborne  Pathogens, OSHA:  29 CFR,  Part  1910.1030.

8.  Occupational  Exposure  to  Hazardous  Chemicals  in  Laboratories, 29 CFR, Part 1910.1450.

9.  Connecticut  General  Statute,  Section  19a-31a    Biolevel-three Laboratories.

10.  Connecticut  General  Statute,  Section  31-40v.  Safety and Health Committees.

11.  Control  of  Communicable  Diseases  Manual,  Abraham S. Benenson, Editor, 16th Edition

An official report of the American Public Health Association.,  1995.  577 pages.

12.  American  Biological  Safety  Association:  www.absa.org    -   Biosafety  Levels.

.

Page  9.