CONNECTICUT  DEPARTMENT  OF  PUBLIC  HEALTH

DIVISION  OF  HEALTH  SYSTEMS  REGULATION

410 Capitol Ave., M.S.# 12HSR, Hartford, CT 06134;  860-509-7400

LABORATORY  BIOSAFETY   INSPECTION  CHECKLIST.         Rev.  6/10/2005

Laboratory: ___________________________________    Registration No.: __________

Address:     _______________________________________________________________

Director:     ____________________________________         Room No.: ____________

Inspector:   _______________________________    Date of Inspection: ____________

Agents Used:	  Bacterial	  Viral 		  Parasitic	  Fungal

Transmissible spongiform encephalopathies	  Rickettsial

Biosafety Level:    2                         3

A. Standard Microbiological Practices: Biosafety Level 2 & 3 Yes No NA

Access to the laboratory is limited or restricted at the discretion of the laboratory director when experiments are in progress.

Persons wash their hands after they handle viable materials, after removing gloves, and before leaving the laboratory.

3. Eating, drinking, smoking, handling contact lenses, and applying cosmetics are not permitted in the work areas. Food is stored outside the work area in cabinets or refrigerators designated for this purpose only.

4. Mouth pipetting is prohibited; mechanical pipetting devices are used.

5. Policies for the safe handling of sharps are instituted.

6. All procedures are performed carefully to minimize the creation of splashes or aerosols.

Work surfaces are decontaminated on completion of work or at the end of the day and after any spill or splash of viable material with disinfectants that are effective against the agents of concern.

All cultures, stocks, and other regulated wastes are decontaminated before disposal by an approved decontamination method such as autoclaving. Materials to be decontaminated outside of the immediate laboratory are placed in a durable, leakproof container and closed for transport from the laboratory. Materials to be decontaminated off-site from the facility are packaged in accordance with applicable local, state, and federal regulations, before removal from the facility.

9. An insect and rodent control program is in effect.

B. Special Practices BSL- 2 Yes No NA

Access to the laboratory is limited or restricted by the laboratory director when work with infectious agents is in progress. In general, persons who are at increased risk of acquiring infection, or for whom infection may have serious consequences, are not allowed in the laboratory or animal rooms. For example, persons who are immunocompromised or immunosuppressed may be at increased risk of acquiring infections. The laboratory director has the final responsibility for assessing each circumstance and determining who may enter or work in the laboratory or animal room.

Intended use:  This inspection checklist is intended to be used for evaluating  the biosafety standards utilized in microbiology laboratories.  A performance based approach must be taken when evaluating compliance with the requirements and recommendations contained in this checklist.  Laboratories may adopt new technologies and innovative techniques to meet the intent of the CDC/NIH BMBL Guidelines, provided that adequate safety  precautions are enforced to protect public health by ensuring containment of  microorganisms.               Page  1.

B. Special Practices BSL- 2 Yes No NA

The laboratory director establishes policies and procedures whereby only persons who have been advised of the potential hazards and meet specific entry requirements (e.g., immunization) may enter the laboratory.

A biohazard sign must be posted on the entrance to the laboratory when etiologic agents are in use. Appropriate information to be posted includes the agent(s) in use, the biosafety level, the required immunizations, the investigator's name and telephone number, any personal protective equipment that must be worn in the laboratory, and any procedures required for exiting the laboratory.

Laboratory personnel receive appropriate immunizations or tests for the agents handled or potentially present in the laboratory (e.g., hepatitis B vaccine or TB skin testing).

When appropriate, considering the agent(s) handled, baseline serum samples for laboratory and other at-risk personnel are collected and stored. Additional serum specimens may be collected periodically, depending on the agents handled or the function of the facility.

Biosafety procedures are incorporated into standard operating procedures or in a biosafety manual adopted or prepared specifically for the laboratory by the laboratory director. Personnel are advised of special hazards and are required to read and follow instructions on practices and procedures.

The laboratory director ensures that laboratory and support personnel receive appropriate training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures. Personnel receive annual updates or additional training as necessary for procedural or policy changes.

A high degree of precaution must always be taken with any contaminated sharp items, including needles and syringes, slides, pipettes, capillary tubes, and scalpels.

Needles and syringes or other sharp instruments should be restricted in the laboratory for use only when there is no alternative, such as parenteral injection, phlebotomy, or aspiration of fluids from laboratory animals and diaphragm bottles. Plasticware should be substituted for glassware whenever possible.

Only needle-locking syringes or disposable syringe-needle units (i.e., needle is integral to the syringe) are used for injection or aspiration of infectious materials. Used disposable needles must not be bent, sheared, broken, recapped, removed from disposable syringes, or otherwise manipulated by hand before disposal; rather, they must be carefully placed in conveniently located puncture-resistant containers used for sharps disposal. Non-disposable sharps must be placed in a hard-walled container for transport to a processing area for decontamination, preferably by autoclaving.

Syringes which re-sheathe the needle, needleless systems, and other safety devices are used when appropriate.

Broken glassware must not be handled directly by hand, but must be removed by mechanical means such as a brush and dustpan, tongs, or forceps. Containers of contaminated needles, sharp equipment, and broken glass are decontaminated before disposal, according to any local, state, or federal regulations.

Cultures, tissues, specimens of body fluids, or potentially infectious wastes are placed in a container with a cover that prevents leakage during collection, handling, processing, storage, transport, or shipping. (See also: DEP Regulation 22a-209-15 Biomedical Waste.)

Page  2.

B. Special Practices BSL- 2 Yes No NA

Laboratory equipment and work surfaces should be decontaminated with an effective disinfectant on a routine basis, after work with infectious materials is finished, and especially after overt spills, splashes, or other contamination by infectious materials. Contaminated equipment must be decontaminated according to any local, state, or federal regulations before it is sent for repair or maintenance or packaged for transport in accordance with applicable local, state, or federal regulations, before removal from the facility.

Spills and accidents that result in overt exposures to infectious materials are immediately reported to the laboratory director. Medical evaluation, surveillance, and treatment are provided as appropriate and written records are maintained.

12. Animals not involved in the work being performed are not permitted in the lab.

C. Safety Equipment (Primary Barriers) BSL-2 Yes No NA

Properly maintained biological safety cabinets, preferably Class II, or other appropriate personal protective equipment or physical containment devices are used whenever:

Procedures with a potential for creating infectious aerosols or splashes are conducted. These may include centrifuging, grinding, blending, vigorous shaking or mixing, sonic disruption, opening containers of infectious materials whose internal pressures may be different from ambient pressures, inoculating animals intranasally, and harvesting infected tissues from animals or embryonate eggs.

High concentrations or large volumes of infectious agents are used. Such materials may be centrifuged in the open laboratory if sealed rotor heads or centrifuge safety cups are used, and if these rotors or safety cups are opened only in a biological safety cabinet.

Face protection (goggles, mask, face shield or other splatter guard) is used for anticipated splashes or sprays of infectious or other hazardous materials to the face when the microorganisms must be manipulated outside the BSC.

Protective laboratory coats, gowns, smocks, or uniforms designated for lab use are worn while in the laboratory. This protective clothing is removed and left in the laboratory before leaving for non-laboratory areas (e.g., cafeteria, library, administrative offices). All protective clothing is either disposed of in the laboratory or laundered by the institution; it should never be taken home by personnel.

Gloves are worn when hands may contact potentially infectious materials, contaminated surfaces or equipment. Wearing two pairs of gloves may be appropriate. Gloves are disposed of when overtly contaminated, and removed when work with infectious materials is completed or when the integrity of the glove is compromised. Disposable gloves are not washed, reused, or used for touching "clean" surfaces (keyboards, telephones, etc.), and they should not be worn outside the lab. Alternatives to powdered latex gloves should be available. Hands are washed following removal of gloves.

D. Laboratory Facilities (Secondary Barriers) BSL-2 Yes No NA

Provide lockable doors for facilities that house restricted agents (as defined in 42 CFR 72.6).

2. Consider locating new laboratories away from public areas.

3. Each laboratory contains a sink for handwashing.

B 6. Note:   If the laboratory adopts  the CDC/NIH BMBL as its biosafety manual,  the laboratory must also:

(a)  maintain  a  list the agents used (inventory), or types of agents it is prepared to handle if encountered in clinical specimens, (i.e. would the lab work-up possible BT agents?), (b)  perform  a risk assessment (identify the Biosafety Level of each agent), and (c) identify the procedures and equipment to be used in handling such agents.

(This information is  recorded on pages 2 and 3 of  the laboratory  registration  application form).       Page  3.

D. Laboratory Facilities (Secondary Barriers) BSL-2 Yes No NA

The laboratory is designed so that it can be easily cleaned. Carpets and rugs in laboratories are inappropriate.

Bench tops are impervious to water and are resistant to moderate heat and the organic solvents, acids, alkalis, and chemicals used to decontaminate the work surfaces and equipment.

Laboratory furniture is capable of supporting anticipated loading and uses. Spaces between benches, cabinets, and equipment are accessible for cleaning. Chairs and other furniture used in laboratory work should be covered with a non-fabric material that can be easily decontaminated.

Install biological safety cabinets in such a manner that fluctuations of the room supply and exhaust air do not cause the biological safety cabinets to operate outside their parameters for containment. Locate biological safety cabinets away from doors, from windows that can be opened, from heavily traveled laboratory areas, and from other potentially disruptive equipment so as to maintain the biological safety cabinets' air flow parameters for containment.

8. An eyewash station is readily available.

Illumination is adequate for all activities, avoiding reflections and glare that could impede vision.

There are no specific ventilation requirements. However, planning of new facilities should consider mechanical ventilation systems that provide an inward flow of air without recirculation to spaces outside of the laboratory. If the laboratory has windows that open to the exterior, they are fitted with fly screens.

B. Special Practices BSL-3 Yes No NA

1. Laboratory doors are kept closed when experiments are in progress.

The laboratory director controls access to the laboratory and restricts access to persons whose presence is required for program or support purposes. Persons who are at increased risk of acquiring infection or for whom infection may have serious consequences are not allowed in the laboratory or animal rooms. For example, persons who are immunocompromised or immunosuppressed may be at risk of acquiring infections. The director has the final responsibility for assessing each circumstance and determining who may enter or work in the laboratory. No minors should be allowed in the laboratory.

The laboratory director establishes policies and procedures whereby only persons who have been advised of the potential biohazard, who meet any specific entry requirements (e.g., immunization), and who comply with all entry and exit procedures, enter the laboratory or animal rooms.

When infectious materials or infected animals are present in the laboratory or containment module, a hazard warning sign, incorporating the universal biohazard symbol, is posted on all laboratory and animal room access doors. The hazard warning sign identifies the agent, lists the name and telephone number of the laboratory director or other responsible person(s), and indicates any special requirements for entering the laboratory, such as the need for immunizations, respirators, or other personal protective measures.

Laboratory personnel receive the appropriate immunizations or tests for the agents handled or potentially present in the laboratory (e.g., hepatitis B vaccine or TB skin testing), and periodic testing as recommended for the agent being handled.

Page  4.

B. Special Practices BSL-3 Yes No NA

Baseline serum samples are collected as appropriate and stored for all laboratory and other at-risk personnel. Additional serum specimens may be periodically collected, depending on the agents handled or the function of the laboratory.

A biosafety manual specific to the laboratory is prepared or adopted by the laboratory director and biosafety precautions are incorporated into standard operating procedures. Personnel are advised of special hazards and are required to read and follow instructions on practices and procedures.

Laboratory and support personnel receive appropriate training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures. Personnel receive annual updates or additional training as necessary for procedural changes.

The laboratory director is responsible for ensuring that, before working with organisms at Biosafety Level 3, all personnel demonstrate proficiency in standard microbiological practices and techniques, and in the practices and operations specific to the laboratory facility. This might include prior experience in handling human pathogens or cell cultures, or a specific training program provided by the laboratory director or other competent scientist proficient in safe microbiological practices and techniques.

A high degree of precaution must always be taken with any contaminated sharp items, including needles and syringes, slides, pipettes, capillary tubes, and scalpels.

Needles and syringes or other sharp instruments should be restricted in the laboratory for use only when there is no alternative, such as parenteral injection, phlebotomy, or aspiration of fluids from laboratory animals and diaphragm bottles. Plasticware should be substituted for glassware whenever possible.

Only needle-locking syringes or disposable syringe-needle units (i.e., needle is integral to the syringe) are used for injection or aspiration of infectious materials. Used disposable needles must not be bent, sheared, broken, recapped, removed from disposable syringes, or otherwise manipulated by hand before disposal; rather, they must be carefully placed in conveniently located puncture-resistant containers used for sharps disposal. Non-disposable sharps must be placed in a hard-walled container for transport to a processing area for decontamination, preferably by autoclaving.

Syringes which re-sheathe the needle, needleless systems, and other safe devices are used when appropriate.

Broken glassware must not be handled directly by hand, but must be removed by mechanical means such as a brush and dustpan, tongs, or forceps. Containers of contaminated needles, sharp equipment, and broken glass should be decontaminated before disposal, and disposed of according to any local, state, or federal regulations.

All open manipulations involving infectious materials are conducted in biological safety cabinets or other physical containment devices within the containment module. No work in open vessels is conducted on the open bench. Clean-up is facilitated by using plastic-backed paper toweling on non-perforated work surfaces within biological safety cabinets.

Laboratory equipment and work surfaces should be decontaminated routinely with an effective disinfectant, after work with infectious materials is finished, and especially after overt spills, splashes, or other contamination with infectious materials.

Page  5.

B. Special Practices BSL-3 Yes No NA

Spills of infectious materials are decontaminated, contained and cleaned up by appropriate professional staff, or others properly trained and equipped to work with concentrated infectious material. Spill procedures are developed and posted.

Contaminated equipment must be decontaminated before removal from the facility for repair or maintenance or packaging for transport, in accordance with applicable local, state, or federal regulations.

Cultures, tissues, specimens of body fluids, or wastes are placed in a container that prevents leakage during collection, handling, processing, storage, transport, or shipping.

All potentially contaminated waste materials (e.g., gloves, lab coats, etc.) from laboratories are decontaminated before disposal or reuse.

15. Spills and accidents that result in overt or potential exposures to infectious materials are immediately reported to the laboratory director. Appropriate medical evaluation, surveillance, and treatment are provided and written records are maintained.

Animals and plants not related to the work being conducted are not permitted in the laboratory.

C. Safety Equipment (Primary Barriers) (BSL-3) Yes No NA

Protective laboratory clothing such as solid-front or wrap-around gowns, scrub suits, or coveralls are worn by workers when in the laboratory. Protective clothing is not worn outside the laboratory. Reusable clothing is decontaminated before being laundered. Clothing is changed when overtly contaminated.

Gloves must be worn when handling infectious materials, infected animals, and when handling contaminated equipment.

Frequent changing of gloves accompanied by hand washing is recommended. Disposable gloves are not reused.

All manipulations of infectious materials, necropsy of infected animals, harvesting of tissues or fluids from infected animals or embryonate eggs , etc., are conducted in a Class II or Class III biological safety cabinet (see Appendix A)(1) .

When a procedure or process cannot be conducted within a biological safety cabinet, then appropriate combinations of personal protective equipment (e.g., respirators, face shields) and physical containment devices (e.g., centrifuge safety cups or sealed rotors) are used.

6. Respiratory and face protection are used when in rooms containing infected animals.

D. Laboratory Facilities (Secondary Barriers) (BSL-3) Yes No NA

The laboratory is separated from areas that are open to unrestricted traffic flow within the building, and access to the laboratory is restricted. Passage through a series of two self-closing doors is the basic requirement for entry into the laboratory from access corridors. Doors are lockable (see Appendix F)(1). A clothes change room may be included in the passageway.

Each laboratory room contains a sink for handwashing. The sink is hands-free or automatically operated and is located near the room exit door.

Page  6.

D. Laboratory Facilities (Secondary Barriers) (BSL-3) Yes No NA

The interior surfaces of walls, floors, and ceilings of areas where BSL-3 agents are handled are constructed for easy cleaning and decontamination. Seams, if present, must be sealed. Walls, ceilings, and floors should be smooth, impermeable to liquids and resistant to the chemicals and disinfectants normally used in the laboratory. Floors should be monolithic and slip-resistant. Consideration should be given to the use of covered floor coverings. Penetrations in floors, walls, and ceiling surfaces are sealed. Openings such as around ducts and the spaces between doors and frames are capable of being sealed to facilitate decontamination.

Bench tops are impervious to water and are resistant to moderate heat and the organic solvents, acids, alkalis, and those chemicals used to decontaminate the work surfaces and equipment.

Laboratory furniture is capable of supporting anticipated loading and uses. Spaces between benches, cabinets, and equipment are accessible for cleaning. Chairs and other furniture used in laboratory work should be covered with a non-fabric material that can be easily decontaminated.

6. All windows in the laboratory are closed and sealed.

A method for decontaminating all laboratory wastes is available in the facility and utilized, preferably within the laboratory (i.e., autoclave, chemical disinfection, incineration, or other approved decontamination method). Consideration should be given to means of decontaminating equipment. If waste is transported out of the laboratory, it should be properly sealed and not transported in public corridors.

Biological safety cabinets are required and are located away from doors, from room supply louvers, and from heavily-traveled laboratory areas.

A ducted exhaust air ventilation system is provided. This system creates directional airflow which draws air into the laboratory from "clean" areas and toward "contaminated" areas. The exhaust air is not recirculated to any other area of the building. Filtration and other treatments of the exhaust air are not required, but may be considered based on site requirements, and specific agent manipulations and use conditions. The outside exhaust must be dispersed away from occupied areas and air intakes, or the exhaust must be HEPA-filtered. Laboratory personnel must verify that the direction of the airflow (into the laboratory) is proper. It is recommended that a visual monitoring device that indicates and confirms directional inward airflow be provided at the laboratory entry. Consideration should be given to installing an HVAC control system to prevent sustained positive pressurization of the laboratory. Audible alarms should be considered to notify personnel of HVAC system failure.

HEPA-filtered exhaust air from a Class II biological safety cabinet can be recirculated into the laboratory if the cabinet is tested and certified at least annually. When exhaust air from Class II safety cabinets is to be discharged to the outside through the building exhaust air system, the cabinets must be connected in a manner that avoids any interference with the air balance of the cabinets or the building exhaust system (e.g., an air gap between the cabinet exhaust and the exhaust duct). When Class III biological safety cabinets are used they should be directly connected to the exhaust system. If the Class III cabinets are connected to the supply system, it is done in a manner that prevents positive pressurization of the cabinets (see Appendix A)(1) .

Page  7.

D. Laboratory Facilities (Secondary Barriers) (BSL-3) Yes No NA

Continuous flow centrifuges or other equipment that may produce aerosols are contained in devices that exhaust air through HEPA filters before discharge into the laboratory. These HEPA systems are tested at least annually. Alternatively, the exhaust from such equipment may be vented to the outside if it is dispersed away from occupied areas and air intakes.

Vacuum lines are protected with liquid disinfectant traps and HEPA filters, or their equivalent. Filters must be replaced as needed. An alternative is to use portable vacuum pumps (also properly protected with traps and filters) .

13. An eyewash station is readily available inside the laboratory.

Illumination is adequate for all activities, avoiding reflections and glare that could impede vision.

15. Biosafety Level 3 facility design and operational procedures must be documented. The facility must be tested for verification that the design and operational parameters have been met prior to operation. Facilities should be re-verified, at least annually, against these procedures as modified by operational experience.

Additional environmental protection (e.g., personnel showers, HEPA filtration of exhaust air, containment of other piped services and the provision of effluent decontamination) should be considered if recommended by the agent summary statement, as determined by risk assessment, the site conditions, or other applicable federal, state, or local regulations.

E. Standard Requirements;

Chemical Safety: Yes No NA

1. The laboratory has a Chemical Hygiene Plan readily available to employees.

2. Employees are provided training to ensure that they are aware of hazards.

3. Written emergency procedures are available in case of spills and accidents.

4. The chemical hygiene officer inspects storage facilities to ensure that incompatible reagents are not stored together.

5. Compressed gas cylinders are secured.

Bloodborne Pathogens: Yes No NA

1. The laboratory has an exposure control plan readily available to employees.

2. Employees are provided training to ensure that they are aware of hazards.

3. The employer provides personal protective equipment, including laboratory coats, latex gloves, and face shields (if applicable) and laundering for laboratory coats.

4. Written emergency procedures are available.

Plumbing: Backflow Prevention Yes No NA

Sink faucets with hose attachments have a backflow preventer, or the hose does not reach below the maximum water height of the sink.

Page  8.

F. Public Health Laboratories; From Reference # 3. Yes No NA

1. Classification of Bioterrorism Network Laboratories: Level A Lab: BSL2 Labs that use Biosafety Cabinets / Rule Out and forward Organisms. Level B Lab: Work at BSL2 – with BT Agents / Confirmation, transport. These include some large clinical labs, State Health Dept. & Large Public Health Labs. Level C Lab: Work at BSL3–Molecular Assays, reference capacity. Public Health Labs with advanced capacity for performing all of the tests of a Level B lab and can selectively perform confirmatory testing, toxigenicity testing or molecular fingerprinting (typing) procedures. Level D Lab: Work at BSL4 - Highest level characterization (Federal, CDC, USAMRIID).

2. Critical Biological Agents: Category A agents: High priority agents- Variola major (smallpox), Bacillus anthracis (anthrax), Yersinia pestis (plague), Francisella tularensis (tularemia), Clostridum botulinum (botulism), Filoviruses (Ebola, Hemorrhagic fever, Marburg hemorrhagic fever). Category B agents: Next highest priority agents include but are not limited to: Coxiella burnetti (Q fever), Brucella species (Brucellosis), Burkholderia mallei (Glanders), Alphaviruses (Venezuelan encephalomyelitis, Eastern and Western encephalomyelitis), Ricin toxin from Ricinus communis (castor beans), Epsilon toxin from Clostridium perfringens, Enterotoxin from Staphylococcus aureus; A subset of B agents include pathogens that are food or water borne. These pathogens include but are not limited to: Salmonella species, Shigella dysenteriae, Escherichia coli O157:H7, Vibrio cholerae, Cryptosporidium parvum. Category C agents: The third highest priority agents include: Nipah virus, Hantaviruses, Tickborne hemorrhagic fever viruses, Tickborne encephalitis virus, Yellow fever virus, Multidrug-resistant Mycobacterium tuberculosis.

Agent					Recommended  Biosafety  Level

Bacterial Agents: Bacillus anthracis					BSL2 with BSL3 Practices &  BSL3

Brucella (B. abortuus, B. melitensis, B. suis)	BSL2 with BSL3 Practices &  BSL3

Burkholderia mallei									 BSL3

Clostridium botulinum toxin			BSL2 with BSL3 Practices &  BSL3

Coxiella burnetii									 BSL3

Franciscella tularensis					BSL2 with BSL3 Practices &  BSL3

Staphylococcus aureus enterotoxin		BSL2 with BSL3 Practices

Yersinia pestis					BSL2 with BSL3 Practices &  BSL3

Viral Agents:	Arboviruses (Eastern & Western Equine 		BSL2 with BSL3 Practices &  BSL3

encephalomyelitis viruses)

Arboviruses (Venezuelan equine 						 BSL3

encephalomyelitis virus, Yellow fever virus)

Arenaviruses (Lassa)					BSL 4

Filoviruses (Ebola, Marburg)			BSL 4

Poxviruses (Variola major)			BSL 4  (Ref. #3, Pg. 76)

Biological  Safety  Cabinets  (BSC):

Class I 75 FPM; Personnel / Env. Protection / Not Product Protection; 100% HEPA filtered exhaust to room

Class  II, Type A 75 fpm opening. HEPA air is 70% Recirculated to inside cabinet / 30 % Exhaust to room

Class II, Type B1  100 fpm; HEPA Air from above, classic design; 30% Recirculated / 70 % Bldg.Exhaust

Class II, Type B2  100 fpm; HEPA air  above;  Radionuclides / Toxic Chemicals;   100 % Bldg. Exhaust

Class II, Type B3  100 fpm opening, front grill, ducted to rear side; 70% Recirculated / 30 % Bldg. exhaust

Class III  fpm is N/A, supply & exhaust is double filtered; Total Containment - Gas Tight - Glove Port –

For BL-3/4; 70% Recirculated / 30 % Bldg. Exhaust. May have dunk tank or pass-through box to autoclave.

Page 9.

Biosafety Level 1 represents a basic level of containment that relies on standard microbiological practices with no special primary or secondary barriers recommended, other than a sink for handwashing. (1)

Biosafety Level 2 practices, equipment, and facility design and construction are applicable to clinical, diagnostic, teaching, and other laboratories in which work is done with a broad spectrum of indigenous moderate-risk agents that are present in the community and associated with human disease of varying severity. Biosafety Level 2 is appropriate when work is done with any human-derived blood, body fluids, tissues, or primary human cell lines where the presence of an infectious agent may be unknown. Laboratory personnel working with human-derived materials should refer to the OSHA Bloodborne Pathogen Standard for specific requirements. (1)

Biosafety Level 3 is applicable to clinical, diagnostic, teaching, research, or production facilities in which work is done with indigenous or exotic agents which may cause serious or potentially lethal disease as a result of exposure by the inhalation route. Laboratory personnel have specific training in handling pathogenic and potentially lethal agents, and are supervised by competent scientists who are experienced in working with these agents. All procedures involving the manipulation of infectious materials are conducted within biological safety cabinets or other physical containment devices, or by personnel wearing appropriate personal protective clothing and equipment. The laboratory has special engineering and design features. It is recognized, however, that some existing facilities may not have all the facility features recommended for Biosafety Level 3 (i.e., double-door access zone and sealed penetrations). In this circumstance, an acceptable level of safety for the conduct of routine procedures, (e.g., diagnostic procedures involving the propagation of an agent for identification, typing, susceptibility testing, etc.), may be achieved in a Biosafety Level 2 facility, providing 1) the exhaust air from the laboratory room is discharged to the outdoors, 2) the ventilation to the laboratory is balanced to provide directional airflow into the room, 3) access to the laboratory is restricted when work is in progress, and 4) the recommended Standard Microbiological Practices, Special Practices, and Safety Equipment for Biosafety Level 3 are rigorously followed. The decision to implement this modification of Biosafety Level 3 recommendations should be made only by the laboratory director.(1)

Universal Precautions are intended to apply only to the transmission of blood-borne pathogens. Standard Precautions cover the transmission of all infectious agents by combining universal precautions and body substance isolation practices. All patient and laboratory specimens are treated as infectious and handled according to “standard precautions”. See: Guideline for Isolation Precautions in Hospitals, Infection Control and Hospital Epidemiology. CDC. 1996;17(1):53-80 and MMWR 1988;37:377-388).

For Public Health, Bioterrorism Network Laboratories:  Risk assessment takes into consideration that BT Agents may be modified by biotechnology and exhibit characteristics different from the known agents.  The biosafety level may be increased  due to increased  virulence, enhanced stability, increased infectivity, drug or vaccine resistance, modified route of     transmission, modified diagnostic characteristics, modified host range infecting increased number of species,  or multiple agents, chimeric agents or combinations of biological and chemical agents.

Biocontainment and safe handling is primary consideration.  Security and chain of custody is also a priority.  Security from unauthorized access includes a single point of separation between authorized personnel and the public or individually secure rooms for each lab employee.   Freezers used to store microbiological agents (Select Agents) must be locked.  Locked storage space meets chain-of-custody requirements.  Chain of custody is a sequential record of each person who has control of a material, including the validity and security of the facility, equipment, test records, and data associated with any observation, collection,  handling, testing and storage of the evidence.

The BSL3 lab has two self-closing doors that are spaced so that they are not both open  simultaneously  during routine entry or egress.   Penetrations through walls, ceilings or floors must be sealed (or be sealable) so that the laboratory rooms can be decontaminated (gas decontamination procedures, e.g with para-formaldehyde).  The ante-room is used as a transition room, where gowns, gloves, respirators, etc.

are put on  and where frequently needed laboratory supplies are stored.   BSL3 labs are not accessible to the general public or to personnel not authorized to enter.  (Reference #3.)                                      Page  10.

Biosafety  Training  Record

Training was provided and the employee(s) has/have the safety skills, knowledge and awareness to work in the following areas (check all that apply):

Biosafety Level:    2               3

Agents:	   Bacterial	  Viral 		  Parasitic	  Fungal

Transmissible spongiform encephalopathies	  Rickettsial

Agents Used: (Specify:  _________________________________________________).

Animals /   Insects (Specify: ________________________________________). Equipment used in combination with infectious agents (check all that apply):

Centrifuge;     Sonicator;     Vortex;     Needles / Sharps;     Autoclave.

Basic Principles:		          YES	NO   NA					YES	NO    NA

Biosafety Levels:			  	      	Risk Assessment:

Route of Transmission:		  	      	Host Susceptibility:

Work Practice Controls: 	  	      	Lab Animal Handling:

Personal Protective Equipment	    	Emergency Response Plan:

Engineering Controls:		  	      	Biological Spill:

Biomedical Waste Disposal:	  	    	Chemical Spill:

Laboratory Ergonomics:	  	    	Accidental Exposure:

Shipping / Transport:					CDC / NIH  BMBL:

Etiological Agents: 		  	    	Standard Practices:

Clinical Specimens: 		  	    	Special Practices:

Specimen Handling /Storage:   	    	Safety Equipment:

Biological Safety Cabinet:				Security:

Principle of Operation:		  	    	Restricted Access:

Using Your BSC:			  	    	Visitor Policy:

Do's  &  Don'ts: 			  	    	Culture Inventory:

Additional Topics Discussed:

Bloodborne Pathogens;    Formaldehyde;                    Chemical Hygiene;

Tuberculosis;                     Hepatitis B Vaccination;     Fire Safety;

Other (Specify): ________________.     Date(s) of  Training: _________________________.

_______________________________________            ___________________________________

Signature of Employee                                                    Signature of Employee

__________________________________	______________________________

Signature of Employee		                                     Signature of Employee

__________________________________	______________________________

Signature of Employee	                                                 Signature of Employee

__________________________________	______________________________

Signature of Employee		                                     Signature of Employee

Training is provided prior to initial assignment of tasks.  Additional training or  updates  are provided when

there is a change in  procedures or risk assessments.

________________________________________

Signature of Director or Supervisor

Page 11.

References

1.  	Biosafety  in  Microbiological  and  Biomedical  Laboratories,   4th Edition,  May 1999.

US Govt. Printing Office.  Tel: 202-512-1800;  HHS Pub. No. (CDC) 93-8395;  Cost:  $12.00

Internet Copy Available Free  At:  www.cdc.gov  and   www.absa.org

- Anthology of Biosafety, III.  Application of Principles,  Jonathan Y. Richmond, Ph.D., Editor

Chapter 1. Risk Assessment for Working with Infectious Agents in the Biological Laboratory,

Richard C. Knudsen, Ph.D.;   Available at  www.absa.org

3.  Anthology of Biosafety, IV. Issues in Public Health, Edited by Jonathan Y. Richmond, Ph.D.

Chapter 14, Biosafety in Public Health Laboratories.  Available at  www.absa.org

4.   Designing Modern Microbiological / Biomedical Laboratory: Lab Design, Process & Technology.

Jonathan Y. Richmond, Ph.D., Editor,   Available at:  www.apha.org

5.   American Type Culture Collection at:  www.atcc.org    ( Click on:  "Search Our Catalogs",

to obtain  Biosafety Levels ).

6.   Guidelines for  Preventing  the Transmission of  Tuberculosis  in  Health  Care  Facilities,

CDC MMWR, 45 (RR-8) June 7,  1996.  Available at www.cdc.gov

7.   Procedures for the  Handling  and  Transport  of  Diagnostic  Specimens and  Etiologic  Agents.

Third Edition, NCCLS Approved Standard H5-A3  NCCLS   Tel: (215) 525-2435, www.nccls.org

- Guideline for Hand Hygiene in Health-Care Settings.   Recommendations of the Healthcare Infection

Control Practices Advisory Committee and HIPAC / SHEA / APIC / IDSA Hand Hygiene Task Force.

Available at www.cdc.gov   and   www.absa.org    MMWR Oct. 25, 2002  51(RR16); 1-44

9.   Immunization of Health - Care Workers;  Recommendations of the Advisory Committee on

Immunization Practices (ACIP) and the Hospital Infection Control Practices Committee (HICPAC)

MMWR  Dec. 26, 1997,  Vol. 46 /  No. RR-18;  Available at :  www.cdc.gov

10.  Cumitech 29.  Laboratory Safety in Clinical Microbiology.  July 1996.

American Society for Microbiology;   www.asmusa.org

11.  Occupational Exposure to Bloodborne Pathogens, OSHA: 29 CFR, Part 1910.1030.

12.  Protection of Laboratory Workers From Occupationally Acquired Infections: Approved Guideline;

NCCLS Document M29;   Clinical Laboratory Standards Institute (CLSI)

Formerly:  National Committee for Clinical Laboratory Standards.;

Available from:  NCCLS, 940 West Valley Rd. suite 1400, Wayne, PA  19087-1898     www.nccls.org

13.  APIC  Bioterrorism Readiness Plan – Template – www.cdc.gov/ncidod/diseases/bioterr.htm

14.   Guidelines  For  Research  Involving  Recombinant  DNA  Molecules  (NIH  Guidelines),

Available at  www.absa.org

15.   NIEHS  Health  and  Safety  Guide  to  Laboratory  Ergonomics.

Division of Safety, National Institutes of Health;   Available at  www.nih.org   Ergonomics (search)

16.   Occupational  Exposure  to  Hazardous  Chemicals  in  Laboratories, 29 CFR, Part 1910.1450

17.  Control  of  Communicable  Diseases  Manual,  Abraham S. Benenson, Editor, 16th Edition

An official report of the American Public Health Association.,  Available at www.apha.org

- BOCA National Plumbing Code, P-1507.9 -1507.12  Backflow protection.

Page 12.